## 128 Slice CT
MEDICAL DEVICE SPECIFICATION (Including Information on the following where relevant/appropriate, but not limited to)
128 Slice CT Scanner
Version no. : |  | 1
Date: |  | 22/12/2015
Done by : (name / institution) |  | HCT/NHSRC
NAME AND CODING
GMDN name |  | 128 Slice CT Scanner
GMDN code |  | 37618
GENERAL
1 | USE
1.1 | Clinical purpose | Computed Tomography scanners are used for a wide variety of diagnostic procedures, including spine and head injuries, lesions, and abdominal and pelvic malignancies; to examine the cerebral ventricles, the chest wall, and the large blood vessels; and to assess musculoskeletal degeneration.
1.2 | Used by clinical department/ward | Radiology Department
TECHNICAL
2 | TECHNICAL CHARACTERISTICS
2.1 | Technical characteristics (specific to this type of device) | 1. Gantry a) Should incorporate low Voltage Slip Rings b) Minimum scan time for a 360º rotation should be less than or equal to 0.35 sec. (350 mili sec.) c) Should have minimum tilt of 30 degrees on either side and remote tilt should be available as standard d) Gantry should be provided with user control panels on either side for positioning of the patient e) The sub millimeter slice @0.63 mm or less in 64 row 128 acquisitions should be available. The system should be in position to perform 128 slices / rotation for general, cardiac and vascular applications f) Should have 3D positioning laser lights g) The scan FOV in acquisition mode be at least 200 mm to 500 mm with intermediate steps for scanning different anatomies h) Gantry aperture should be at least 70 cm. in diameter i) Integrated Display Panel - Gantry front showing current scan parameters such as kV, mA, ECG trace etc. for easy set up for ECG gated studies. 2. X-Ray Generator a) Should be compact and in-built in the gantry b) Should be high frequency having at least 100 kW output or more c) The mA range available should be between 20 to 800 or more, with increment steps of not more than 10 mA. d) Tube Voltage: 80-140 kV 3. X-Ray Tube a) The X-ray tube should be dual focus with heat storage capacity of 8 MHU or more, with effective storage of at least 25 MHU. b) Peak heat dissipation rate of anode should be at least 1600 KHU/min c) X-ray tube cooler unit should be inside the gantry d) Focal spots, and type of X-ray tube should be specified as per IEC Recommendations. e) Filter and beam limiting devices should be quoted as standard.
2.1 | Technical characteristics (specific to this type of device) | 4. Detectors a) These should be of solid state type b) 128 Slice acquisition per rotation should be possible with the detectors, in 0.63 mm mode. c) The system should have at least 64 ‘physical rows’ of the detectors.Number of elements in each row should be specified d) The Z-axis coverage of at least 40 mm / rotation should be possible for standard and cardiac scans e) Fan-angle of X-rays and the geometry should be specified f) Detectors should not require frequent calibration 5. Patient Table a) Should have minimum weight bearing capacity of 200 kilograms b) The minimum table top height should not be more than 35 cms from floor level for easy transport of trauma patients c) Table top width to be at least 42 cms d) The range of metal free scannable range should be at least 160 cm. e) The vertical range (max. Ht. – min. Ht.) 55 cm f) Remote controlled UP / DOWN and FWD / BWD movement. g) Pitch to be freely selectable in automatic / manual mode : 0.15 – 1.5 h) Reproducing accuracy of the Table : 1mm
2.1 | Technical characteristics (specific to this type of device) | 6. Spiral CT capabilities: a) Minimum slice thickness should be 0.63 mm or less and maximum 10 mm or more. b) Pitch factor (volume pitch): Variable between 0.5 to 1.5 or more and should be user selectable. c) Spiral length: 150cm or more. d) Single continuous 'spiral-on time' should be minimum 100 seconds or more. e) Bolus triggered spiral acquisition should be possible. f) True isotropic volume acquisition and sub-millimeter resolution of at-least 0.4 mm for all body applications. 7. Topogram:- a) Length and width: specify range. b) Scan times: specify range c) Should be possible to interrupt acquisition manually once the desired anatomy is obtained. 8. Data acquisition system: a) System should have minimum 64 rows of detector capable of generating 128 slices through latest flying focal spot technology or equivalent. b) Mention minimum acquired slice thickness in Axial & Helical mode after reconstruction. c) Acquisition of cardiac images with ECG gating (prospective & retrospective) should be possible d) Step and shoot technique during cardiac scanning for dose reduction, or a similar alternative technology should be available. 9. Image Evaluation Tools: a) Parallel evaluation of multiple ROI in circle, irregular and polygonal forms. b) Statistical evaluation for area/volume, S.D., Mean, Min/Max and histogram. c) Distance and angle measurement, freely selectable positioning of co-ordinate system, grid and image annotation.
2.1 | Technical characteristics (specific to this type of device) | 10. Latest Iterative Reconstruction Technique: a) ASIR-V/ iDose4 Premium / SAFIRE or latest available with the manufacturer to be quoted as standard. b) Model-based Iterative reconstruction technology VEO/ IMR/ ADMIRE or equivalent for all imaging protocols including hardware and software. c) Low dose protocols for pediatric and infant scanning.
2.1 | Technical characteristics (specific to this type of device) | 11. Image Reconstruction: a) Real time reconstruction speed: 20 images per second or more at 512 x 512 matrix. b) Display matrix: 1024 x 1024 or more. c) Reconstructed slice thickness range should be less than one mm (<1) to 10mm. d) Patient's radiation Dose must be displayed on monitior and Imaging Films. 12. Image Quality: a) The high contrast resolution be more than 20 lp/mm in all routine scans, including spiral and axial mode b) The low contrast resolution should not be more than 3 mm at 0.5%
2.2 | User's interface | Patient Communication System: An integrated intercom and automated patient instruction system (API) should be provided.
2.3 | Software and/or Workstation | Workstations: A client server architecture based solution (Intellispace Portal 6/ Dexus-AW server 2/ Syngo Via 30A or equivalent.) with minimum concurrent 24,000 slices rendering capacity, with storage of minimum 1TB having following client hardware specifications-Workstation: Z820 or equivalent CPU, dual quad core processor, 16 GB RAM, 1TB hard drive, DVD Writing with clinical grade monitor of minimum 2 MP.A reputed Anti-Virus Solution for Server should be in place. The Server should be with minimum three user (Three Hardware’s) facility Fully DICOM 3.0 Compliant and PACS Interface ready. The workstation should have following processing tools/software’s Available as standard: • Multi planar reconstruction(MPR) , • Minimum and Maximum intensity projection • 3D Volume rendering , • 3D SSD (Shaded Surface Display). • Advance Vessel Analysis with plaque visualization,• Auto Bone Removal. • Volume measurement, • Lung Nodule analysis. • Liver lesion analysis. • Colonography. • Perfusion CT. • Image Fusion of CT, MR & PET Data • Neuro DSA. • Coronary tree analysis: automated 3D processing of coronary arteries, calcium scoring, stent analysis, LV analysis • Multi-modality automatic tumour tracking & Automatic measurements in RECIST, WHO, Volume & Choi criteria calculation. • Virtual endoscopy.
3 | PHYSICAL CHARACTERISTICS
3.1 | Dimensions (metric) | NA
3.2 | Weight (lbs, kg) | NA
3.3 | Configuration | NA
3.4 | Noise (in dBA) | NA
3.5 | Heat dissipation | Suitable Heat Sink/Cooling Mechanism to be provided.
3.6 | Mobility, portability | Stationary Installation
4 | ENERGY SOURCE (electricity, UPS, solar, gas, water, CO2 ....)
4.1 | Power Requirements | NA
4.2 | Battery operated | No
4.3 | Tolerance (to variations, shutdowns) | NA
4.4 | Protection | Suitable Servo Stablizer to be provided
4.5 | Power consumption | NA
5 | ACCESSORIES, SPARE PARTS, CONSUMABLES
5.1 | Accessories (mandatory, standard, optional); | a)Dry Chemistry Laser Imager (dpi 500 or more) of a reputed make : Integrated with main console and workstation b) Color Laser Printer (High Resolution) for color coaded images C) UPS with half hour ‘back-up’ to run entire CT system, Workstations and Laser Imager (should be 160 kVA or more, d) Dual – Head Pressure injector of reputed make (100 syringes) e) 160 KVA Silent DG Set with AMF panel f) Two LED based view boxes with adjustable illumination to view 3 films of 14” x 17” in each view box. g) Thyroid Collars -2 No. h) Gonadal Shields- 2 each for male and female(Total 4) i) Lead Apron Hanger with 2 light weight Lead Aprons
BIDDING / PROCUREMENT TERMS / DONATION REQUIREMENTS
6 | ENVIRONMENTAL AND DEPARTMENTAL CONSIDERATONS
6.1 | Atmosphere / Ambiance (air conditioning, humidity, dust ...) | Operating condition: Dehumidifiers to be provided to maintain humidity between 30 to 70 % Air conditioning of the whole complex to maintain temperature range between 15 to 25 Celsius.
6.2 | User's care, Cleaning, Disinfection & Sterility issues | 1)Disinfection: Parts of the Device that are designed to come into contact with the patient or the operator should either be capable of easy disinfection or be protected by a single use/disposable cover. 2)Sterilization not required.
7 | STANDARDS AND SAFETY
7.1 | Certificates (pre-market, sanitary, ..); Performance and safety standards (specific to the device type);Local and/or international | 1.Should be FDA/ European CE 2.Manufacturer and Supplier should have ISO 13485 certification for quality standards. 3.Electrical safety conforms to the standards for electrical safety IEC 60601-1-General requirements. 4.Shall meet internationally recognised standard for Electromagnetic Compatibility(EMI/EMC) for electromedical equipment: 61326-1. 5. Equipment should be AERB type approved
7.2 | Local and/or international | Manufacturer should have ISO 13485 certificate for quality standard.
8 | TRAINING AND INSTALLATION
8.1 | Pre-installation requirements: nature, values, quality, tolerance | Turnkey Project only space to be provided, bidder shall provide all the fixtures and furniture along with interiors as per AERB norms.
8.2 | Requirements for sign-off | Certificate of calibration and inspection of parts from the manufacturer
8.3 | Training of staff (medical, paramedical, technicians) | 1)Training of users on operation and basic maintenance on site for a minimum period of 4 weeks. 2)Advanced maintenance tasks required shall be documented
9 | WARRANTY AND MAINTENANCE
9.1 | Warranty | 5 years
9.2 | Maintenance tasks | CMC 5 years 2 PM Visits Annually. All Breakdown calls to be attended within 24 hrs of registartion.
9.3 | Service contract clauses, including prices | The spare price list of all spares and accessories (including minor) required for maintenance and repairs in future after guarantee / warranty period should be attached;
10 | DOCUMENTATION
10.1 | Operating manuals, service manuals, other manuals | Should provide 2 sets(hardcopy and soft-copy) of:- 1)User, technical and maintenance manuals to be supplied in english/hindi language along with machine diagrams; 2)List of equipment and procedures required for local calibration and routine maintenance; 3)Service and operation manuals (original and copy) to be provided; 4)Advanced maintenance tasks documentation; 5)Certificate of calibration and inspection
10.2 | Other accompanying documents | List of essential spares and accessories, with their part numbers and cost;
11 | NOTES
11.1 | Service Support Contact details (Hierarchy Wise; including a toll free/landline number) | Contact details of manufacturer, supplier and local service agent to be provided; Any Contract (AMC/CMC/add-hoc) to be declared by the manufacturer;
11.2 | Recommendations or warnings | Any warning signs would be adequately displayed