## MRI
MEDICAL DEVICE SPECIFICATION (Including Information on the following where relevant/appropriate, but not limited to)
Full Body MRI System
Version no. : |  | 1
Date: |  | 24/08/2015
Done by : (name / institution) |  | HCT/NHSRC
NAME AND CODING
GMDN name |  | Full Body MRI System Super Conducting Magnet
GMDN code |  | 37654
GENERAL
1 | USE
1.1 | Clinical purpose | MRI is primarily used to identify diseases of the central nervous system, brain, and spine and to detect musculoskeletal disorders.It is also used to view cartilage, tendons, and ligaments. MRI can also be used to image the eyes and the sinuses. MRI can be used to help diagnose infectious diseases; to detect metastatic liver disease; to display heart-wall structure; to stage prostate,bladder, and uterine cancer.MRI can also be used as a functional imaging tool.
1.2 | Used by clinical department/ward | Radiology Department
TECHNICAL
2 | TECHNICAL CHARACTERISTICS
2.1 | Technical characteristics (specific to this type of device) | 1. MAGNET a. Whole Body 1.5Tesla Magnetic Resonance Imaging System optimized for higher performance in Whole Body and Vascular examinations with superconducting magnet, high performance gradients and digital Radio Frequency System. b. 1.5T active shielded super conductive magnet should be short and non- claustrophobic. c. It should have at least 70 cm patient bore with flared opening. d. Magnet length should be less than 200cm. e. Homogeneity of magnet should be less than 3.5 ppm over 45cm DSV f. The magnet should be well ventilated and illuminated with built in 2 way intercom for communication with patient. g. It should have a built in cryo-cooler such that helium consumption does not exceed 0.01 lit/ hour. h. Emergency Rundown Control at both operator console room and Gantary Room is a must. i. Fringe Field 0.5 Gauss line radius is essential. j. Front Panel of gantry should display table and patient position. 2. SHIM SYSTEM a. High performance, highly stable shim system with global and localized automated shimming for high homogeneity magnetic field for imaging and spectroscopy. b. Auto shim should be available to shim the magnet with patient in position. 3. GRADIENT SYSTEM a. Actively shielded Gradient system b. The gradient should be actively shielded with each axis having independently a slew rate of at least 200 T/m/s and a peak amplitude of 33mT/m. c. The system should have efficient and adequate Eddy current compensation d. Effective cooling system for gradient coil and power supply e. Duty Cycle- 100% the gradient power amplifier. f. Usable over 45 cm of FOV in all directions.
2.1 | Technical characteristics (specific to this type of device) | 4. RF SYSTEM a. A fully digital RF system capable of transmitting power of at least 15kw. b. It should also have at least 32 independent RF receiver channels with each having bandwidth of 1 MHz or more along with necessary hardware to support quadrature ICP array/Matrix coils. The highest receiver channels available with the vendor should be quoted. c. It should support Parallel acquisition techniques with a factor of up to 4 in 2D. d. Should allow remote selection of coils and / or coil elements. 5. PATIENT TABLE a. The table should be fully motorized, computer controlled table movements in vertical and horizontal directions. b. A CCTV system with colour LCD display to observe the patient should be provided: Moving table angiography should be possible. c. There should be a hand held alarm for patients d. Light Localizer for patient positioning. e. Physiological signals display like ECG, Pulse and SPO2 f. Patient load bearing capacity, minimum 200 Kg. 6. MEASUREMENT SYSTEM a. Largest Field of View should be at least 45 cm in all axis. b. The measurement matrix should be from 128x128 to 1024x1024. c. Minimum 2D slice thickness mm should be equal to or less than 0.5 d. Minimum 3D slice thickness mm should be equal to or less than 0.1
2.1 | Technical characteristics (specific to this type of device) | 7. COIL SYSTEM a. The main body coil integrated to the magnet must be Quadrature / CP. In addition to this following coils should be provided: b. Multichannel Head coils with at least 12 channel for high resolution brain imaging. c. Neuro-vascular Coil with 16 or more channels or Head / Neck Coil combined, capable of high resolution neuro-vascular imaging d. 18 Channel Spine Array/Matrix Coils for thoracic and lumbar spine imaging. e. Body Array/Matrix coil with at least 40 cm z axis coverage for imaging of abdomen, angiograms and heart with 32 channel. f. Suitable Cardiac Coil g. Dedicated 8 channel extremity coil. h. Bilateral Breast Coil with at least 4 channel with fully functional spectroscopy. i. Dedicated Shoulder Coil with atleast 8 channels j. Dedicated Knee Coil with atleast 8 channels k. General purpose flexible coil with small and large size l. Coil Storage Cart from manufacturer. m. The system should continuously monitor the RF coils used during scanning to detect failure modes. RF coils should not require either set up time or coil tuning; Multi coil connection for up to 2 or more coils simultaneous scanning without patient repositioning. 8. APPLICATION SEQUENCES a. The system should have basic sequences package with Spin Echo, InversionRecovery, Turbo Spin Echo with high turbo factor of 256 or more, Gradient Echo with ETL of 255 or more, FLAIR. b. Single slice, multiple single slice, multiple slice, multiple stacks, radial stacks and 3D acquisitions for all applications. c. Single and Multi shot EPI imaging techniques with ETL factor of 255 or more d. Fat suppression for high quality images both STIR and SPIR. e. The system should acquire motion artifact free images in T2 studies of brain in restless patients
2.1 | Technical characteristics (specific to this type of device) | f. Dynamic study for pre and post contrast scans and time intensity studies g. MR angio Imaging: Should have 20/30 TOF, 20/30 PC , MTS and TONE, ceMRA, Facilities for Accelerated time resolved vascular imaging with applications like Treats/Tracks/Tricks sequences. h. Fat and water excitation package. Diffusion Weighted Imaging, with at least b value of 5000 or more. i. Bolus chasing with automatic and manual triggering from fluro mode to 3D acquisition mode with moving table facility. j. Non contrast enhanced peripheral angiography for arterial flow with Native/Trance/lnhance sequences k. Whole body screening imaging studies for metastasis I. High resolution Abdominal and Liver imaging in breathold and free breathing modes with respirator triggered volume acquisitions m. The system should have basic and advanced MRCP packages including free breathing and 3D techniques. n. The system should have facility for flow quantification of CSF, vessel flow and hepatobiliary system.
2.1 | Technical characteristics (specific to this type of device) | o. The system should have the Hydrogen, Single Voxel spectroscopy,Multivoxel,Multislice & Multiangle 2D, 3D Spectroscopy and Chemical shift imaging in 2D/3D. The complete processing/post-processing software including color metabolite maps should be available on main console. Complete prostate spectroscopy hardware and applications should be provided. p. Advanced Cardiac Applications: VCG gating, Morphology/wall motion; Cine perfusion imaging; Myocardial viability imaging; Arrhythmia rejection techniques, Advanced Cardiac Ventricular Measurement Analysis; Cine Cardiac Tagging Techniques; Coronary artery techniques; real time interactive imaging, 20/30 fast field echo/balanced/steady state techniques and evaluation package on workstation q. Advanced Breast imaging Package. r. Perfusion imaging of brain (including PASL and CASL) s. Susceptibility weighted imaging with phase information. t. Multi Direction DWl and DTI with minimum of 32 directions(Complete package including quantification and tractography software). Prospective motion correction enabled software preferred. u. High resolution imaging for inner ear
2.1 | Technical characteristics (specific to this type of device) | 10. SAFETY FEATURES a. The magnet system should include an Emergency Ramp Down unit (ERDU) for fast reduction of the magnetic field with Ramp Down time below 3 minutes b. The magnet should have .quench bands that contain the fringe fields to a specified value in the event of a magnet quench c. Real time SAR calculation should be performed by software to ensure that RF power levels comply with regulatory guidelines and are displayed on each image d. The system shall have manual override of the motor drive for quick removal of the patients from the magnet bore e. Temperature sensor (built in) for magnet refrigeration efficiency must be provided
2.2 | User's interface | a. The main Host computer should have a 19 inches or more high resolution LCD TFT color monitor with 1024 x 1024 matrix display b. The system should have image storage capacity of 100 GB for at least 2,00,000 images in 256x256 matrix. c. The reconstruction speed should be at least 1300 or more for full FOV 256 matrix. d. The main console should have facility for music system for patient in the magnet room. The system should have DVD / CD / flash drive archiving facility. The system should be provided with auto DVD writer. e. Two way intercom system for patient communication. f. MRI System should be DICOM ready in all parameters with no additional requirement of licence for connectivity to any PACS/HIS and Radiotherapy treatment planning system.
2.3 | Software and/or Workstation | a. A workstation with same user interface as of main console is required with the availability of all necessary software including: i. Basic post-processing software including MIP, MPR, surface reconstruction and volume rendering technique. ii. Advanced post-processing offered applications perfusion quantification, advanced diffusion and DTI, processing of 20/30 CSI data, with color metabolite mapping, quantification of CSF flow data, vascular analysis package. b. It should have at least 19 inch LCD TFT color monitor, with hard disk of at least 120 GB for at least 250,000 image storage in 256 matrix, and 4 GB RAM capacity or more, with self playing DVD/CD archiving facility. c. Seperate viewing station should be provided. d. The workstation should enable printing in laser film camera and color printers
3 | PHYSICAL CHARACTERISTICS
3.1 | Dimensions (metric) | NA
3.2 | Weight (lbs, kg) | NA
3.3 | Configuration | NA
3.4 | Noise (in dBA) | Maximum 120 dBA
3.5 | Heat dissipation | Heat Dissipation: Should maintain nominal Temp and the heat should be disbursed through a cooling mechanism with less than 1˚ c change during scan
3.6 | Mobility, portability | Stationary Installation
4 | ENERGY SOURCE (electricity, UPS, solar, gas, water, CO2 ....)
4.1 | Power Requirements | ??????
4.2 | Battery operated | No
4.3 | Tolerance (to variations, shutdowns) | ??????
4.4 | Protection | ?????
4.5 | Power consumption | maximum 40 KW
5 | ACCESSORIES, SPARE PARTS, CONSUMABLES
5.1 | Accessories (mandatory, standard, optional); Spare parts (main ones); Consumables / reagents (open, closed system) | a. Dual Head MRI Compatible Pressure Injector with 100 sets of syringes. b. Water Chiller for Cold Head I Gradients.. c. 2 Non-ferromagnetic patient transfer trolley of international make should be provided. d. Fire Fighting System, Detectors and 6 Fire Extinguishers. e. Hand held metal detectors and two mental detector doors to be installed at the entrance point as will be intimated. f. Closed circuit CCD camera g. Phantoms for image quality audits. h. MRI compatible Anaesthesia machine i.Suction and O2 pipeline and manifold to be provided inside the RF enclosure. j. Suitable RF Enclosure k. UPS for entire system for backup of 30 minutes.
BIDDING / PROCUREMENT TERMS / DONATION REQUIREMENTS
6 | ENVIRONMENTAL AND DEPARTMENTAL CONSIDERATONS
6.1 | Atmosphere / Ambiance (air conditioning, humidity, dust ...) | 1)Operating condition: Capable of operating continuously in ambient temperature of 5 to 50 deg C and relative humidity of 15 to 80% .
6.2 | User's care, Cleaning, Disinfection & Sterility issues | 1)Disinfection: Parts of the Device that are designed to come into contact with the patient or the operator should either be capable of easy disinfection or be protected by a single use/disposable cover. 2)Sterilization not required.
7 | STANDARDS AND SAFETY
7.1 | Certificates (pre-market, sanitary, ..); Performance and safety standards (specific to the device type);Local and/or international | 1.Should be FDA/ European CE 2.Manufacturer and Supplier should have ISO 13485 certification for quality standards. 3.Electrical safety conforms to the standards for electrical safety IEC 60601-1-General requirements. 4.Shall meet internationally recognised standard for Electromagnetic Compatibility(EMI/EMC) for electromedical equipment: 61326-1.
7.2 | Local and/or international | Manufacturer / supplier should have ISO 13485 certificate for quality standard.
8 | TRAINING AND INSTALLATION
8.1 | Pre-installation requirements: nature, values, quality, tolerance | Turnkey Project only space to be provided.?????
8.2 | Requirements for sign-off | Certificate of calibration and inspection of parts from the manufacturer
8.3 | Training of staff (medical, paramedical, technicians) | 1)Training of users on operation and basic maintenance on site for a minimum period of 4 weeks. 2)Advanced maintenance tasks required shall be documented
9 | WARRANTY AND MAINTENANCE
9.1 | Warranty | 3 years
9.2 | Maintenance tasks | CMC 5 years 2 PM Visits Annually. All Breakdown calls to be attended within 24 hrs of registartion.
9.3 | Service contract clauses, including prices | The spare price list of all spares and accessories (including minor) required for maintenance and repairs in future after guarantee / warranty period should be attached;
10 | DOCUMENTATION
10.1 | Operating manuals, service manuals, other manuals | Should provide 2 sets(hardcopy and soft-copy) of:- 1)User, technical and maintenance manuals to be supplied in english/hindi language along with machine diagrams; 2)List of equipment and procedures required for local calibration and routine maintenance; 3)Service and operation manuals (original and copy) to be provided; 4)Advanced maintenance tasks documentation; 5)Certificate of calibration and inspection
10.2 | Other accompanying documents | List of essential spares and accessories, with their part numbers and cost;
11 | NOTES
11.1 | Service Support Contact details (Hierarchy Wise; including a toll free/landline number) | Contact details of manufacturer, supplier and local service agent to be provided; Any Contract (AMC/CMC/add-hoc) to be declared by the manufacturer;
11.2 | Recommendations or warnings | Any warning signs would be adequately displayed