## Slide 1: National summit on best practices & innovations in public healthcare systems 3-6 July 2013 SKICC Srinagar TEMPERATURE DATA LOGGERS: Piloting in DISTRICT Kurukshtra (Haryana)
Dr Suresh Dalpath
DDCH/SEPIO
Government of Haryana
AFFORDABLE HEALTH TECHNOLOGIES (PHFI) & IMMUNIZATION TECHNICAL SUPPORT UNIT (ITSU)- MOHFW, GOI

## Slide 2
Universal Immunization Programme (UIP) was launched in 1985 and has completed 2.5 decades of successful implementation
Each Year, >12 crore children and 4 crore pregnant mothers in India are targeted for vaccination with various antigens including 20 lakh children in Haryana
As per CES-2009, 71.6 % children in Haryana & 61% children Nationally are fully immunized
India has introduced Pentavalent vaccine in 8 states including Haryana and many other potential vaccines are currently available.
Background

## Slide 3
Under national immunization programme BCG, DPT, Hepatitis B, Polio, Measles and TT vaccines are universally provided
In addition, JE and Pentavalent vaccines are provided in limited geographical areas
All vaccines under Immunization Programme are temperature sensitive ( both high and low) and need to be stored at +2 to + 8 deg C or -20 deg C depending upon type of vaccine and storage level
Background…2

## Slide 4
Cold chain needs to be maintained from Manufacture up to Session site
Variety of active (ILR, DF, WIC, WIF) and passive cold chain equipment (Cold boxes, Vaccine Carriers) have been provided in the system to store and maintain vaccines at recommended temperature
  All vaccines vials are provided with VVM ( Vaccine vial monitor) for monitoring heat exposure to vaccines.
Violation of temperature range leads to
  loss of potency of vaccine
  Probability of an Adverse Effect
Background…3
Cardinal principal of immunization: “Don’t Harm”

## Slide 5
As per guidelines, temperature of each cold chain equipment should be recorded minimum twice a day including Sundays & holidays
Cold chain handlers need to record temp., monitoring by MO, DIO, CMO
Manual Recording
In temperature Log books using thermometers
Automated/Semi-Automated
WIC/WIF are provided with 7 or 30 day temperature graph recorder and hooter
Few WIC/WIF also equipped with wireless data loggers
Temperature Recording: Current Practices

## Slide 6
Manual Recording: Compliance issues
Failure to record temperature twice daily
Non-recording of temperature on Holidays
Details of power failure rarely documented
Differential interpretation of thermometer readings by cold chain handlers
Not a real time data and historic records cannot be verified
Inadequate sensitization of health staff
Automated/Semi-Automated
Larger setup required/high investment cost
Dependent of electricity supply and/or internet availability
Non-availability of usable parts ( ink of graph recorder)
Limitation of Current Temperature Monitoring
ALL temperature monitoring is currently localized with No dynamic system to monitor cold chain performance

## Slide 7
Cold Chain Maintenance:
Studies and Field Visits Observations

## Slide 8
Unicef: Cold chain evaluation report 2008
≈

## Slide 9
Chart: 
| Supervisory visits | ILRs & DFs away from walls and other equipment | ILRs & DFs away from direct exposure | ILRs & DFs connected through Stabilizers | Temperature Log Books for all ILRs & DFs | Twice daily monitoring of temperature | Record of power failures/cuts | Record of Defrosting ILRs & DFs | Periodic checks of Temperature Log Books
Round 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Observations of RAPID Visit
≈
≈
≈
≈
≈

## Slide 10
Some observations during field visits

## Slide 11
Some observations during field visits

## Slide 12
Some observations during field visits cntd..
How can we be sure that vaccines given to beneficiaries in immunization programme is of good quality?

## Slide 13
Use of SIM enabled temperature data logger:
An innovative yet simple technique

## Slide 14
Real time visibility of Temperature at cold chain equipment
Use of information for timely response
Accountable and vigilant system with proper trail.
Functional in limited resource setting
OBJECTIVE OF USE
The Ultimate objective is
“ Safe and Potent Vaccine For Each Child”

## Slide 15
Control
Buttons
Monitor
Screen
A special electronic device designed by PHFI
Developed to monitor temperature of cold chain equipment
Has an In-built SIM to send out real time temperature readings and alerts at calibrated time interval
Temperature Data Logger: The Device

## Slide 16
Deep Freezer
Box Type
ILR Type
Walk in Cooler
Walk in Freezer
Ice Lined Refrigerator
Deep Freezer
Cold Boxes
Vaccine carriers
Temperature Data Logger: Where it can be used
WIC

## Slide 17
Sensor are placed with vaccine and attached to a vaccine vial
Wire of sensor is very thin and can enter the equipment without effecting closing and opening of door
The temperature sensor is attached to vial itself for accurate measurements
Temperature Data Logger: How it cn be used

## Slide 18
Temperature Data Logger: In ILR & DF

## Slide 19
2-3 sensors inside the WIC & WIF are placed to measure temperature of different corners
Temperature Data Logger: In WIC & WIF

## Slide 20
SENSOR MODULE
The temperature sensing module constantly monitors the temperature of the CC equipment. When the temperature goes beyond recommended range, the SMS module is activated.
SMS MODULE
The SMS module sends out an SMS to a pre-fed mobile number about the temperature status of the CC Equipmenty
Temperature Data Logger: How the system works

## Slide 21
The Temprature logger Pilot : Haryana
Haryana initiated the pilot project throughout the compete Cold Chain system i.e. from State Vaccine Store upto last cold chain point in Kurukshetra district
Supported by PHFI/ITSU-GOI
The Pilot was launched in April 2013.
31 Temperature Loggers installed,
  2 In State Vaccine Store at Panchkula,
  4 at regional vaccine store KKR
  25 in PHC/ CHC cold chain points of KKR dist.
Initial troubleshooting and stabilization of devices going on.
Contingency and Response Plan developed for responding to temperature failures

## Slide 22
The Temprature Logger Pilot : Haryana
State Vaccine store

## Slide 23
The Temprature logger Pilot : Haryana Contd..
State Level orientation on response plan competed
Conducted an inspection drill exercise for all machines initially reporting faulty temperature
Caliberation issues resolves and troubleshooting done
Identified nodal person for each cold chain point
Nodal person will receive an SMS alert whenever there is a temperature violation
District level orientation scheduled in July 2013.

## Slide 24
Temperature Data Logger: Monitoring Portal
Temperature Loggers installed are mapped on a central monitoring portal
Monitoring Portal have details about the logger device like
  GPS Location,
  Centre In charge,
  Contact Details,
  History of carrier temperature data,
  graphical representation etc..
Centre In Charges are given access to the portal so that they can monitor the carrier temperature.

## Slide 25
Temperature Data Logger: Monitoring Portal

## Slide 26
Temperature Data Logger: Monitoring Portal

## Slide 27
Monitoring & Feedback Mechanism
Intervention
Monitoring

## Slide 28
Response Plan in event of Temp Failure
Temp more than recommended
Temp less than recommended

## Slide 29
SMS Alerts
SMS Alert Protocol |  |  | 
Level | Equipment | Condition | Alert to be Sent to
PHC/Other Sub District Cold Chain Points | ILR | 2 Degree and below 8 Degree and above | Cold Chain Handler MO In-charge
 | If no correction at local level for 12 hour (or > 2 outlier readings in 24 hrs), alert for individual equipment will be sent to- |  | District Cold Chain Technician Dist. CC Handler DIO
District Cold Chain Points | ILR | 2 Degree and below 8 Degree and above | District Cold Chain Technician Dist. CC Handler DIO
 | DF | > ( - 15) Degree | 
Regional Cold Chain Point | WIC | 2 Degree and below 8 Degree and above | Reg. CC Handler DIO Dist. CC Technician State CCO State CC Tech SIO
 | WIF | > - 15 Degree | 
State Cold Chain Point | WIC | 2 Degree and below 8 Degree and above | State CCO State CC Tech SIO
 | WIF | > - 15 Degree | 

## Slide 30
Conclusion
Temprature monitoring is very important to ensure delivery of safe vaccine to beneficiary
System has started recording the temperature for the cold chain equipment and transmitting in real time to the website
It helped in sensitising the health staff for prompt action.
Initial response has been very encouraging
Scaling up appears to be very much feasible as well as utterly desirable.

## Slide 31
Way forward
This Pilot has been recently launched with installation of Temperature monitoring devices
A full scale monitoring and response will start in July 2013 in pilot areas.
Contingency plan is to be implemented with mechanism of prompt response to cold chain break down & temperature failure
Improved system performance from the pilot is expected only after 6 -12 months with scope of scale up to other areas.

## Slide 32
Thanks
Team Haryana